| QUICK REFERENCE GUIDE FOR POCT |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| KIT NAME | TEST | TESTING TEMP | TESTING TIME | SAMPLE TYPE | QC MATERIAL | QC FREQUENCY | STORAGE TEMP |
| OSOM HCG COMBO TEST | PREG | 15-30◦C | 3 MINUTES | RANDOM URINE (first morning is best) | QUANTRIMETRIX DROPPER URINALYSIS CONTROL | MONTHLY OR WHEN NEW BOX IS OPENED | 15-30◦C |
| SIEMENS MULTISTIX 10 SG | URINE DIP | 22-26◦C | 60s to 2 MINUTES | RANDOM URINE (first morning is best) | QUANTRIMETRIX DROPPER URINALYSIS CONTROL | DAILY OR WHEN A PATIENT IS TESTED | 22-26◦C |
| CLEARVIEW MONO | MONO | 15-30◦C | 5 MINUTES | EDTA WHOLE BLOOD (LAV) | SUPPLIED IN KIT | MONTHLY OR WHEN NEW BOX IS OPENED | 2-30◦C |
| QUICKVUE IN-LINE STREP A TEST | RAPID STREP | 15-30◦C | 5 MINUTES | THROAT SWAB SUPPLIED WITH KIT OR QUICKVUE EXTERNAL QC | SUPPLIED IN KIT | MONTHLY OR WHEN NEW BOX IS OPENED | 15-30◦C |
| GLUCOMETER | GLUCOSE | 15-30◦C | VARIES | FINGERSTICK | ACCU-CHECK AVIVA QC | DAILY OR WHEN A PATIENT IS TESTED | 2-30◦C (TEST STRIPS) |
| LIFE SIGN MI TROPONIN I | TROPONIN-I | 15-30◦C | 15 MINUTES | HEPARIN WHOLE BLOOD (MINT) | RUN BY LAB | NEW BOX QC'D BY LAB | 2-30◦C |
| ISTAT | chem8+ | 15-30◦C | APPROX 2 MIN | HEPARIN WHOLE BLOOD (MINT) | ISTAT CONTROLS LEVELS 1,2 AND 3 | MONTHLY OR WHEN NEW BOX IS OPENED | 2-8◦C |
|  | CG4+ | 15-30◦C | APPROX 2 MIN | HEPARIN WHOLE BLOOD (MINT) | ISTAT CONTROLS LEVELS 1,2 AND 3 | MONTHLY OR WHEN NEW BOX IS OPENED | 2-8◦C |
|  | cTnl TROPONIN I | 15-30◦C | APPROX 10 MIN | HEPARIN WHOLE BLOOD (MINT) | ISTAT cTnl CONTROLS LEVELS 1,2 AND 3 | MONTHLY OR WHEN NEW BOX IS OPENED | 2-8◦C |